## Slide 1: Acoustic Measurement Systems 101
Bob Smith
Principal Scientist, Acoustic Systems Physio-Control

## Slide 2: Agenda
Overview of an Acoustic Measurement System
System Calibration
Microphones
Measurement System Comparisons
  Dedicated Hardware vs. Computer + SW
Audio Interfaces
Example real world use
Demonstrations

## Slide 3
Input Overview

## Slide 4
Output Overview

## Slide 5
1 Channel System

## Slide 6
2 Channel System

## Slide 7
3 Channel System

## Slide 8: System Calibration
Microphone mV/Pa
Gain staging between mic and digital domain
Digital domain 0dBFS representation
How many mV / Pa at microphone = ?? dBFS in digital domain
Mic calibrator (level only)
  Bruel & Kjaer 4231 Type 1 NIST traceable ~ $2K
  Landtek ND9 – not traceable ~ $120

## Slide 9: System Calibration
If the mv/Pa is known for a measurement mic the following can be used in place of a mic calibrator to get close to calibrated SPL levels:
  Signal generator (even an iPod will work)
  AC voltmeter with a mV scale
  Set the sig gen for 1 KHz (or use 1 KHz wavefile)
  Connect sig gen to voltmeter and adjust output such that the voltmeter reads the same mV as the mic mV/Pa
  Now connect the sig gen output between Pins 2 & 3 of the audio I/O mic preamp input (P48 off!) and adjust the preamp gain for the dBFS representation desired

## Slide 10: Gain Staging: -0.05 dBFS no clipping

## Slide 11: Gain Staging: 0 dBFS clipping

## Slide 12: Gain Staging: -2 dBFS inter-stage clipping

## Slide 13: System Cal: 94 dBSPL = -16 dBFS

## Slide 14: Measurement Microphones
Bruel & Kjaer, ACO Pacific, Microtech Gefell, DPA
Josephson, Earthworks, LinearX
Beyerdynamic MM1, Audix TM1
Dayton EMM-6, Nady CM-100, The B word
DIY Panasonic wm61a

## Slide 15: Measurement Microphones
Frequency Response Accuracy
Off axis performance (how omni is omni?)
Self Noise
Dynamic Range
Maximum SPL

## Slide 16: DPA 4006 Polar Response

## Slide 17
 |  |  | dB A | mV/Pa | mV/Pa | Hz - Hz | dB | dB | $ | 
mfg | model | type | self noise | sensitivity | measured | freq resp | tol +/- | max SPL | price | notes
ACO+Josephson | 7046+C617 | ext | 12 | 67 | 70.9 | 2-20K | 2.0 | 145 | 2,000 | 
Audix | TM-1 | pre | 28 | 6.5 |  | 20-25K | 2 | 130 | 300 | 
Beyerdynamic | MM-1 | pre | 28 | 15 | 14.9 | 50-16K | 1.5 | 128 | 200 | 
dbx | M2 | pre | NS | 7 | 6.46 | 20-20K | NS | NS | 100 | 
Behringer | ECM-8000 | pre | NS | NS | 10 ~ 14 | 15-20K | 2.5 | NS | 50 | unbal
Nady | CM-100 | pre | NS | 10 | 9.18 | 20-20K | NS | 120 | 50 | unbal
Dayton | EMM-6 | pre | 24 ? | 10 | 8.99 | 18-20K | NS | 127 | 50 | unbal
Earthworks | M-30 | pre | 27 | 8 |  | 9-27K | 1.0 | 150 | 650 | 
Earthworks | M-50 | pre | 22 | 30 |  | 3-50K | 1.0 | 142 | 1,300 | 
DPA | 4003 | pre | 15 | 40 |  | 10-20K | 2.0 | 135 | 1,925 | 
DPA | 4006 | pre | 15 | 10 |  | 10-20K | 2.0 | 124 | 1,925 | 
Sennheiser | MKH-20 | ext | 10 | 25 |  | 12-20K | 2.0 | 135 | 1,200 | 
Schoeps | CMC6+MK2 | ext | 11 | 15 |  | 20-20K | 0.5 | 130 | 1,800 | 
Josephson | C550H | pre | 26 | 10 | 9.85 | 20-20K | 2.0 | 135 | 400 | 
Josephson | 606+KA11N | ext | 15 | 15 |  | 20-20K | 2.0 | 140 | 1,280 | NA

## Slide 18: Microphone Calibration
Reciprocity – primary standard method
  IEC 61094-2, edition 2
Free Field Comparison
  IEC 61094-5, edition 1
Pistonphone
  typically 250 Hz, 114 or 124 dBSPL
Electronic Sound Calibrator
  typically 1000 Hz, 94, 104, 114 or 124 dBSPL

## Slide 19: Reference Mic 1

## Slide 20: Reference Mic 2

## Slide 21: Free Field Comparison Calibration

## Slide 22: Measurement Mics Compared

## Slide 23: Measurement Mics Compared

## Slide 24: ACO 7046-28439

## Slide 25: Josephson C550H-5938

## Slide 26: Beyerdynamic MM1-2267

## Slide 27: Dayton EMM6-2372

## Slide 28: Nady CM-100

## Slide 29: dbx M2

## Slide 30: The B company ECM8000 3 mics

## Slide 31: Measurement Systems
Bruel & Kjaer, NT Instruments, Meyer SIM, Audio Precision, Rohde & Schwarz
CLIO, LinearX, DRA MLSSA
EASERA, Smaart, Praxis
Omnimic, TrueRTA, ETF5
Room EQ Wizard, Speaker Workshop

## Slide 32: Dedicated HW vs. Computer + SW
Purpose Specific Imbedded OS instead of Win7, Lion or Linix
System crashes, BSOD occurrence minimized
Dedicated controls for rapid feature access
Minimal or no system integration
System calibration
Can be more $$$

## Slide 33: Rohde & Schwarz UPV

## Slide 34: Audio Precision SYS-2722

## Slide 35: Measurement Systems
Bruel & Kjaer, NT Instruments, Meyer SIM, Audio Precision, Rohde & Schwarz
  Purpose built imbedded OS / hybrid with host OS
  Turnkey
  High Accuracy
  Calibrated
  Comprehensive
  $5K to $50K

## Slide 36: Measurement Systems
CLIO, LinearX, DRA MLSSA
  Hybrid embedded OS / computer host OS
  Minor integration
  High Accuracy
  Calibrated
  Comprehensive
  $2K to $10K

## Slide 37: Measurement Systems
EASERA, Smaart, Praxis
  Host OS
  Moderate integration
  Accuracy depends on quality of audio interface
  Can be calibrated
  EASERA comprehensive, Smaart more focused
  $800 to $1K except EASERA
  $800 to $5K EASERA depending on configuration

## Slide 38: Measurement Systems
Omnimic, TrueRTA, ETF5
  Host OS
  More complex integration except Omnimic
  Accuracy depends on quality of audio interface
  Can be calibrated
  Omnimic is a unique system
  Useful feature sets
  Free to $300

## Slide 39: Measurement Systems
Room EQ Wizard, Speaker Workshop
  Host OS
  DIY integration
  Accuracy depends on quality of audio interface
  Can be calibrated
  Useful feature sets
  Free

## Slide 40: Audio Interfaces
Computer Interface
  Firewire 400 & 800
  USB 1.1, 2, 3
  Computer internal bus
  OS drivers

## Slide 41: Audio Interfaces - Examples
RME Fireface UC $1300
MOTU UltraLite-mk3 $550 (FW + USB)
M-Audio Fast Track Pro $200
Lexicon I-Onix U22 $250
Focusrite Scarlett 2i2 $150
Mackie Onyx Blackjack $150
Tascam US-100 $100

## Slide 42: RME Fireface UC (USB)

## Slide 43: M-Audio Fast Track Pro

## Slide 44
1 Channel System Example

## Slide 45
2 Channel System Example

## Slide 46: Useful Integration Tools
dpc latency checker
  deferred procedure calls timed
Microsoft rattv3
  time spent in ISR, deferred procedure calls and timer dpc
Centrance ASIO latency checker

## Slide 47: Real World Measurements
Speaker frequency response
Measurement system comparison
Microphone comparison
Acoustic door performance

## Slide 49: Genelec 8030a – ACO 7046 mic

## Slide 50: Genelec 8030a – ACO 7046 mic

## Slide 51: Genelec 8030a – Dayton EMM6 mic

## Slide 52: Genelec 8030a – 7046 - EMM6 mics

## Slide 53: Acoustic Door Gasket Sound Leak

## Slide 54: Acoustic Door Gasket Seal Improved

## Slide 55: Demonstrations
Meyer SIM
SmaartLive v7
Speaker measurement – nominal
Speaker measurement – contaminated
Delay measurement
Omnimic – time permitting
Room EQ Wizard – time permitting

## Slide 56: Links to Mfgs mentioned
Measurement Microphones:
http://www.bksv.com/Products/TransducersConditioning/AcousticTransducers/Microphones.aspx
http://www.acopacific.com/micdetal.html
http://www.acopacific.com/chart.html
http://www.microtechgefell.de/index.php/en/measuring-equipment/sound--noise
http://www.dpamicrophones.com/en/products.aspx?c=Catalog&category=234
http://www.josephson.com/pdf/srs5.pdf
http://www.earthworksaudio.com/our-microphones/m-series/
http://www.linearx.com/products/microphones/m51/M51_1.htm
http://north-america.beyerdynamic.com/shop/ct/conference-technology/product-line/microphones/special-microphones/mm-1.html
http://www.audixusa.com/docs/products/TM1.shtml
http://www.daytonaudio.com/index.php/emm-6-electret-measurement-microphone.html
http://www.nady.com/cm100.html
http://www.behringer.com/EN/Products/ECM8000.aspx

## Slide 57: Links to Mfgs mentioned
DIY wm61a measurement mic:
http://www.johncon.com/john/wm61a/
http://www.linkwitzlab.com/sys_test.htm
http://www.speakerbuilder.net/web_files/Articles/diymic/diymicmain.htm
Microphone Calibration Service:
http://www.scantekinc.com/
Microphone Field Calibrators (SPL level):
http://www.bksv.com/products/telecomaudiosolutions/electroacousticsaccessories/soundcalibratortype4231.aspx
http://01640c0.netsolhost.com/fieldcali.htm
http://www.acopacific.com/splcali.html
http://www.ebay.com/sch/i.html?_from=R40&_trksid=m570.l2736&_nkw=nd9+calibrator
Audio Interfaces:
http://www.rme-audio.de/en_products_fireface_uc.php
http://www.motu.com/products/motuaudio/ultralite-mk3
http://www.m-audio.com/products/en_us/FastTrackPro.html
http://www.focusrite.com/products/audio_interfaces/scarlett_2i2/
http://www.mackie.com/products/onyxblackjack/
http://tascam.com/product/us-100/

## Slide 58: Links to Mfgs mentioned
Acoustic Measurement Systems:
http://www.bksv.com/
http://www.nti-audio.com/
http://www.meyersound.com/
http://ap.com/
http://www.rohde-schwarz.us/en/products/test_and_measurement/audio/audio_products/
http://www.audiomatica.com/clio10/home1.htm
http://www.linearx.com/
http://www.mlssa.com/
http://www.renkus-heinz.com/easera/index.html
http://www.rationalacoustics.com/
http://www.libinst.com/
http://www.parts-express.com/pe/showdetl.cfm?Partnumber=390-790
http://www.trueaudio.com/
http://www.etfacoustic.com/
http://www.hometheatershack.com/roomeq/
http://www.audua.com/
http://www.purebits.com/
http://www.sigview.com/

## Slide 59: Links to Useful Integration Tools
Useful Integration Tools:
http://www.thesycon.de/deu/latency_check.shtml
http://msdn.microsoft.com/en-us/windows/hardware/gg487354
http://msdn.microsoft.com/en-us/windows/hardware/gg487334
http://centrance.com/downloads/ltu/